Medi SPICE Development

| Fergal Mc Caffery Regulated Software Research Group, Dundalk Institute of Technology, Dundalk, Ireland. Fergal.mccaffery@dkit.ie | Alec Dorling InterSPICE Ltd. Cambridge, England. alec.dorling@interspice.uk.com |
| --- | --- |

Please send proofs to Fergal Mc Caffery at the following address,

Regional Development Centre, Dundalk Institute of Technology, Dundalk,

Co. Louth, Ireland, Fergal.mccaffery@dkit.ie

Tel +353 42 9370522, Fax +353 42 9331163

Keywords:

Medi SPICE, Software Process Improvement (SPI), ISO/IEC 15504, FDA, Medical Device Directive, Medical Device Software Compliance, IEC 62304.

Abstract

This paper outlines the development of a software process assessment and improvement model (Medi SPICE) for the medical device industry. The paper details how medical device regulations may be satisfied by extending relevant processes and practices from ISO/IEC 15504-5. The paper also describes the proposed phases of delivery for Medi SPICE.

Medi SPICE will define a a Process Reference Model (PRM), Process Assessment Model (PAM) and Organizational Maturity Model (OMM). The Medi SPICE Process Assessment Model will be used to perform ISO/IEC 15504 conformant assessments of the software process capability of medical device suppliers in accordance with the requirements of ISO/IEC 15504-2: 2003. The Medi SPICE Process Assessment Model will be based on ISO/IEC 15504-5: 2006 extended to provide coverage of additional software development practices that are required to achieve regulatory compliance within the medical device industry. The processes will be defined within an Organizational Maturity Model conformant with ISO/IEC TR 15504-7:2008.

1. Introduction

Medical device companies must produce a design history file detailing the software components and processes undertaken in the development of their medical devices. Due to the safety-critical nature of medical device software it is important that highly effective software development practices are in place within medical device companies.

Medical device companies who market within the USA must ensure that they comply with medical device regulations as governed by the FDA (FDA - Food and Drug Administration) [1,2,3,4]. The medical device companies must be able to produce sufficient evidence to support compliance in this area.  To this end, the Center for Devices and Radiological Health (CDRH) has published guidance papers for industry and medical device staff which include risk based activities to be performed during software validation [2], pre-market submission [3] and when using off-the-shelf software in a medical device [4].  Although the CDRH guidance documents provide information on which software activities should be performed, including risk based activities; they do not enforce any specific method for performing these activities.

This paper highlights the need for a software process assessment and improvement model within the medical device industry (Medi SPICE) and draws upon the call for specific software development standards to be developed for the medical device industry so that companies adhering to such standards will have a more streamlined pathway towards FDA compliance.

Medi SPICE has the aim of minimising the volume of software documentation content within the premarket submission to the FDA for audit and to provide global harmonization (with consistent guidance provided for all medical device software manufacture) [5]. The results of a Medi SPICE process assessment may be used to provide objective evidence of the state of a medical device suppliers software practices in relation to the regulatory requirements of the industry, and at the same time identify areas for process improvement. The results of assessments may also be used as a criterion for supplier selection. The authors believe that, with the publication of the Medi SPICE Process Reference Model, Process Assessment Model, and Organizational Maturity Model more specific guidance will be available for the basis of process design and assessment in the medical device industry.

We describe the early stage development of Medi SPICE based upon applicable processes from the ISO/IEC 15504-5 Process Assessment Model [6]. We also illustrate high-level mappings that have been performed between medical device regulations and ISO/IEC 15504-5 processes. The approach for delivering Medi SPICE by the Medi SPICE development team is illustrated in figure 1. The model is flexible in that relevant elements may be adopted to provide the most significant benefit to the business.

Extend ISO/IEC 15504-5 with new goals and practices

ISO/IEC 15504-5

Regulatory bodies

Process

Reference

Model

Process Assessment Model

Test in the medical device industry

Figure 1. Medi SPICE Development

Medi SPICE will consist of an extended set of ISO/IEC 15504-5 processes. However in the initial phase the focus will be on the set of ISO/IEC 15504-5 processes that are deemed most applicable to the regulatory requirements of the medical device industry.

2. Medi SPICE Development

To perform this research project, the team plan to deliver a PRM and a PAM which contain processes providing the comprehensive coverage of all the FDA and European Council guidelines [7], and associated standards (e.g. ISO 14971[8], IEC 60601-1-4 [9], IEC 62304 [10], TIR 32 [11], GAMP [12, 13]) for the complete software development lifecycle. As safety is a primary issue for medical device software, the PRM and PAM will incorporate: the safety integrity levels and the safety life cycle from the international standard for the “functional safety” of electrical, electronic, and programmable electronic equipment (IEC 61508) [14];

the safety processes that are present in +SAFE [15] - these will be a foundation for ISO/IEC 15504-10 Safety Extensions.. ISO/IEC 15504-10 will define  additional processes and guidance to support the use of the exemplar process assessment models for systems and software (ISO/IEC 15504 Parts 5 and 6) when applied to the assessment of safety related systems in order to make consistent judgement regarding process capability and/or improvement priorities.

An overall objective will be to propose a conformity assessment scheme to support first, second or third party assessment results that may be recognised by the regulatory bodies. The PRM and PAM of the Medi SPICE assessment standard will be derived from the ISO/IEC 15504-5 processes as they are all applicable to the development of safety-critical medical device software. As the IEC 62304 standard contains the medical device software lifecycle processes that have to be adhered to in order to achieve medical device regulatory compliance a key objective will be to provide coverage of all processes that are either included in or referenced from IEC 62304.

3. Project Outline

Medi SPICE will consist of a defined set of software processes that will reference base practices which when utilised will assist medical device software development organisations to fulfil the regulatory guidelines and standards of the medical device industry.  Medi SPICE will cover the complete medical device software development and maintenance lifecycle.

Medi SPICE will also provide guidance to assessing the software engineering process capability of a medical device software development organization conformant with the ISO/IEC 15504-2 [16] requirements. Medi SPICE will integrate the defined IEC 62304 processes with the relevant ISO/IEC 15504-5 processes to enable FDA guidelines to be fulfilled. Additionally, the safety processes from +SAFE and various other medical device related standards will be integrated into the relevant ISO/IEC 15504-5 processes e.g. ISO 14971 into the risk management process. The FDA has defined the following eleven software development areas:

- Level of Concern

- Software Description

- Device Hazard and Risk Analysis

- Software Requirements Specification

- Architecture Design

- Design Specifications

- Requirements Traceability Analysis

- Development

- Validation, Verification and Testing

- Revision Level History

- Unresolved Anomalies

The IEC 62304 Medical device software – Software lifecycle processes standard defines a number of process requirements that should be adhered to, these are as follows:

Quality Management System,

Software Safety Classification,

Software Development Process, which includes:

- Software Development Planning

- Software Requirements Analysis

- Software Architectural Design

- Software Detailed Design

- Software Unit Implementation & Verification

- Software Integration & Integration Testing

- Software System Testing

- Software Release

Software Maintenance Process, which includes:

- Establish Software Maintenance Plan

- Problem & Modification Analysis

- Modification Implementation

Risk Management Process, which includes:

- Analysis of software contributing to hazardous situations

- Risk Control measures

- Verification of risk control measures

- Risk management of software changes

Software Configuration Management Process, including:

- Configuration Identification

- Change Control

- Configuration status accounting

Software Problem Resolution Process, which includes:

- Prepare problem reports

- Investigate the Problem

- Advise Relevant Parties

- Use change control process

- Maintain Records

- Analyse problems for trends

- Verify software problem resolution

- Test documentation contents

Medi SPICE is being developed to instill effective practices into the software development processes of medical device companies. The goal is to improve the effectiveness and efficiency of software practices used by medical device companies through investigating the mapping between relevant ISO/IEC 15504-5 processes, the eleven FDA software development areas, the five process area groups defined in IEC 62304 and extended safety processes that will be defined in ISO/IEC 15504-10.

All ISO/IEC 15504-5 processes are applicable to the development of safety-critical medical device software and are candidate processes to be included in Medi SPICE. However, in order to progress Medi SPICE a number of pilot processes will initially be developed. The pilot set of processes will then be expanded until Medi SPICE contains a complete set of processes that will fulfil all the associated regulatory standards and guidelines. Our approach will be to agree upon an initial set of processes that will satisfy the eleven software development areas defined by the FDA, the European Council, the +SAFE processes, and the main process groups (i.e. Software Development, Software Maintenance. Software Risk Management, Software Configuration Management, Software Problem Resolution) that are defined in IEC 62304.

The Medi SPICE PRM will contain a set of process descriptions with; process name, purpose, and outcomes for each of the processes. In addition the Medi SPICE PAM will contain for each process: the set of base practices required to accomplish the process purpose and fulfill the process outcomes; sample output work products; characteristics associated with each work product.

Table 1, illustrates the result of performing a high-level mapping of requested FDA pre-market submission areas and IEC 62304 processes against existing ISO/IEC 15504-5 processes. From performing this mapping it is evident that there is not an exact one-to-one mapping from an FDA area to both an IEC 62304 process and a ISO/IEC 15504-5 process and that overlaps exist. Table 1 has grouped together processes where the most significant overlaps occur, however it should be noted that other less significant overlaps will also occur across groups.  The table includes the 11 software areas that are defined by the FDA, these are mapped against 14 processes that are defined in IEC 62304 plus the 2 + SAFE processes, which are in turn mapped against 23 ISO/IEC 15504-5 processes and 3 additional safety related processes that will be included in ISO/IEC 15504 Part 10 - i.e. Safety management, Safety Engineering & the Selection and qualification of software tools and libraries. It should be noted that whilst other ISO/IEC 15504-5 processes are relevant for developing medical device software, the 23 ISO/IEC 15504-5 processes listed are the most relevant (from the set of 48 processes). These processes will be extended with medical device specific content to include safety integrity levels and to satisfy the associated medical device requirements included in both the FDA guidelines and ISO/IEC 62304.

We initially considered developing Medi SPICE to solely focus upon the medical device regulations and to use the IEC 62304 processes as the foundation. However, our aim is for medical device companies not only to develop software that will adhere to the regulatory guidelines but also to encourage the adoption of medical device software development processes that will lead to the development of safer and more reliable medical device software. We feel that this will be achieved through following established software engineering practices such as those documented in ISO/IEC 15504-5.

Table 1. Mapping FDA software areas and IEC 62304 processes against ISO/IEC 15504-5

| FDA Areas | Associated IEC 62304 Processes | Associated ISO/IEC 15504-5 Processes |
| --- | --- | --- |
| Risk and Safety Management Group |
| Device Hazard and Risk Analysis Level of Concern +SAFE – A Safety extension to CMMI CMMI Process Areas Safety Management Safety Engineering | Risk Management Software Safety Classification | A1. Risk Management A2. Safety Management (Part 10 Safety Extensions) A3. Safety Engineering (Part 10 Safety Extensions) |
| Requirements Group |
| 3. Software Requirements Specification 4. Requirements Traceability Analysis | Software Requirements Analysis | B.1 Requirements Elicitation A.4 Software Requirements Analysis B.2 System Requirements Analysis |
| Development Group |
| 5. Architecture Design Chart 6. Design Specifications 7. Software Description 8. Software Development Environment Description | Software Development Planning Software Architectural Design Software Detailed Design | A.5 Project Management A.6. Selection and qualification of software tools and libraries (Part 10 Safety Extensions) B.3 System Architectural Design A.7 Software Design B.4 Documentation |
| Testing and Integration Group |
| 9. Validation, Verification and Testing | Software Unit Implementation and Verification Software Integration and Integration Testing Software System Testing | A.8 Software Construction A.9 Software Integration A.10 Software Testing B.5 System Integration B.6 System Testing A.11 Verification A.12 Validation |
| Supporting Processes |
| 10. Revision Level History 11. Unresolved Anomalies | 10. Software Release 11. Software Configuration Management 12. Software Maintenance 13. Software Problem Resolution | B.7 Product Release B.8 Product Acceptance Support B.9 Software Installation A.13 Configuration Management A.14 Problem Resolution Management A.15 Change Request Management A.16 Software and System Maintenance |
| Additional Requirements |
|  | 14. Quality Management System | ISO 13485 B.10 Quality Assurance |

From closer inspection of table 1, it may be recognised that 16 of the associated ISO/IEC 15504-5 processes appear in bold and are labelled with an A with the remaining 10 processes (in italics) labelled with a B. The A labelled processes are ISO/IEC 15504-5 processes of which a high proportion of defined base practices will be required either in their current state or in a revised state to satisfy the regulatory demands of the medical device industry.  The labelled B processes are ISO/IEC 15504-5 processes which contain a smaller proportion of defined base practices that will be required either in their current state or in a revised state to satisfy the regulatory demands of the medical device industry.

The 16 labelled A processes are:

Risk Management

Safety Management (to be included in Part 10)

Safety Engineering  (to be included in Part 10)

Selection and qualification of software tools and libraries  (to be included in Part 10)

Software Requirements Analysis

Project Management

Software Design

Software Construction

Software Integration

Software Testing

Verification

Validation

Configuration Management

Problem Resolution Management

Change Request Management

Software and System Maintenance

The 10 labelled B processes are:

Requirements Elicitation

System Requirements Analysis

System Architectural Design

Documentation

System Integration

Product Release

Product Acceptance Support

System Testing

System Installation

Quality Assurance

The composition of the Medi SPICE processes is illustrated in figure 2.

X

Y

Z

X- ISO/IEC 15504-5  Practices that are not mandatory for regulatory compliance.

Y- ISO/IEC 15504-5  Practices that are required for regulatory compliance.

Z- Non-ISO/IEC 15504-5  Practices that are required for regulatory compliance.

Figure 2. Composition of Medi SPICE processes.

Medi SPICE will highlight the additional base practices and outcomes that need to be added to the associated  ISO/IEC 15504-5 processes in order to satisfy medical device regulations (Z), as well as any ISO/IEC 15504-5 outcomes and associated base practices that are not required in order to satisfy medical device regulatory requirements (X).

4. Medi SPICE Delivery Phases

Initial development will be focused on  the 16 A Medi SPICE processes. Table 2, illustrates that Medi SPICE will be delivered in 3 phases. Phase 1 will consist of delivery of the 16 initial  processes. Phase 2, will consist of the delivery of an additional 26 processes and phase 3 will consist of the delivery of all remaining Medi SPICE processes. The mappings between the regulatory guidelines and the relevant ISO/IEC 15504-5 processes will then produce Medi SPICE processes that are currently expected to retain the ISO/IEC 15504-5 process names.

Table 2. Different Phases of Medi SPICE delivery

| Medi SPICE Phases | ISO/IEC 15504-5 Processes to be extended | Resultant Medi SPICE Processes |
| --- | --- | --- |
| Phase 1 | A1. Risk Management A2. Safety Management (Part 10) A3. Safety Engineering (Part 10) A4. Software Requirements Analysis A5. Project Management A6. Selection and qualification of software tools and libraries (Part 10) A.7 Software Design A.8 Software Construction A.9 Software Integration A.10 Software Testing A.11 Verification A.12 Validation A.13 Configuration Management A.14 Problem Resolution Management A.15 Change Request Management A.16 Software and System Maintenance | Risk Management Safety Management Safety Engineering Software Requirements Analysis Project Management Selection and qualification of software tools and libraries Software Design Software Construction Software Integration Software Testing Verification Validation Configuration Management Problem Resolution Management Change Request Management Software and System Maintenance |
| Phase 2 | B1. Requirements Elicitation B2. System Requirements Analysis B3. System Architectural Design B4. Documentation B5. System Integration B6. Product Release B7. Product Acceptance Support B8. System Testing B9. System Installation B10. Quality Assurance | 16 Phase 1 processes + Requirements Elicitation System Requirements Analysis System Architectural Design Documentation System Integration Product Release Product Acceptance Support System Testing System Installation Quality Assurance |
| Phase 3 | 25 Remaining ISO/IEC 15504-5 Processes | 26 (Phase 1& Phase 2) processes + 25 Remaining Processes |

5 Medi SPICE Organizational Maturity

Whilst the main objective is to develop a PRM and PAM that will demonstrate process capability of relevant processes and practices, by defining an Organisational Maturity Model an overall organizational maturity level may be derived from a process capability profile

Table 3, illustrates that a medical device company that implements the Medi SPICE processes will also perform all the base processes that are defined within the ISO/IEC 15504-7 exemplar Organizational Maturity Model [17]. Five of the minimum set of 7 base processes that are required in order to satisfy maturity level 1 are defined as class A processes within Medi SPICE and the other 2 processes are defined as class B processes. Therefore, by satisfying the Medi SPICE requirements for these processes an organisation will have the foundations for achieving a Maturity Level 1 rating.  The additional processes that are included at maturity level 1 are also included within Medi SPICE, 1 of the processes is a class A process and the other 5 are class B processes. Additionally, Medi SPICE contains all the processes that are listed within the minimum set of maturity level 2 processes (i.e. 6 class A and 2 class B processes). However there is no coverage in terms of the processes defined in maturity levels 3, 4 & 5.

Table 3. Mapping Medi SPICE processes to the ISO/IEC 15504-7 Maturity Levels

| Maturity Level 1 Processes (Minimum Set) | Additional Processes |
| --- | --- |
|  | Process | Conditions (Required or Optional) |
| B1. Requirements Elicitation A.4 Software Requirements Analysis A.7 Software Design A.8 Software Construction A.9 Software Integration A.10 Software Testing B.6 Product Release | B2. System Requirements Analysis B.3 System Architectural Design B.5 System Integration B.6 System Testing | Required where development covers system issues and not exclusively software issues. |
|  | B.9 Software Installation | Required where the Organisation Unit is responsible for installing the software product in the customer environment |
|  | A.16 Software and System Maintenance | Required where the Organisation Unit is responsible for ongoing maintenance and evolution of the software and/or system. |
| Maturity Level 2 Processes (Minimum Set) | Additional Processes |
| B.10 Quality Assurance A.11 Verification B.4 Documentation A.13 Configuration Management A.14 Problem Resolution Management A.15 Change Request Management A.5 Project Management A.1 Risk Management | Contract Agreement Supplier Monitoring Customer Acceptance | Required where external or internal suppliers of product components, services or infrastructure are involved in the development projects |
|  | A.12 Validation | Required where confirmation of fitness for use of work products is the responsibility of the Organisation Unit |
|  | Joint Review | Optional where the work in the Organisation Unit involves agreements with stakeholders |
|  | B.8 Product Acceptance Support | Optional where the work in the Organisation Unit involves product acceptance support. |
| Maturity Level 3 Processes | Additional Processes |
| Human Resource Management Training Knowledge Management Infrastructure Process Establishment Process Assessment Process Improvement Organisation Management Quality Management Measurement Audit | Asset Management Reuse Program Management Domain Engineering | Optional if the Organisational Unit has a structured reuse program in force – the three processes are mutually reinforcing. |
| Maturity Level 4 Processes |
| Quantitative Performance Management |
| Maturity Level 5 Processes |
| Quantitative Process Improvement |

4. Medi SPICE Risk Management process

As an example of the work in progress a summary of the risk management process is provided. Risk management is a very important process in relation to the development of safety-critical software for the medical device industry.

The Medi SPICE Risk Management (RM) process seeks to combine the various guidelines and standards within the medical device industry.  It does so in the context of the following regulations: ISO 14971, SW68, IEC 62304, TIR32 and GAMP 4/GAMP 5 and FDA guidance documents. The RM process outlined in this paper was developed following an extensive literature review of standards and CDRH guidance papers which govern the medical device software industry.

This paper describes an integral part of this research by detailing the development of a Medi SPICE RM process that is based upon the ISO/IEC 15504-5 RM process. The Medi SPICE RM process will be an extension of the relevant ISO/IEC 15504-5 RM process which will be specifically tailored to fulfil the RM regulations of the medical device software industry. The RM process may then be adopted by medical device companies to improve their software development practices by providing them with a process that will ensure that their hazard analysis and risk control procedures satisfy the current medical device regulations and guidelines. The Failure Mode and Effects Analysis (FMEA) method of identifying, mitigating and tracking risk and hazard issues will be used within the RM process [18].

The Medi SPICE RM process is developed through mapping regulatory medical device practices against the seven base practices specified in ISO/IEC 15504-5 for RM. These are as follows:

- BP1: Establish risk management scope;

- BP2: Define risk management strategies;

- BP3: Identify risks;

- BP4: Analyse risks;

- BP5: Define and perform risk treatment actions;

- BP6: Monitor risks;

- BP7: Take preventative or corrective action.

Within Medi SPICE each base practice consists of one or more base sub-practices. The following is an example of the “BP1: Establish risk management scope” base practice within Medi SPICE. The aim of this practice is to determine the scope of the RM to be performed. Both ISO/IEC 15504-5 and the medical device regulations specify that the RM scope should be defined. Additionally, the medical standards specify that the strategy should include the life-cycle phases for which the strategy is applicable (see table 4). In the following table a new base practice that was added is represented in bold italics.

Table 4: Medi SPICE RM Base Practices for Establishing RM scope

| Base Practice No: | Description | Specified in ISO/IEC 15504-5? | Specified In the Medical device regulations? |
| --- | --- | --- | --- |
| Man.5.BP1 | Determine the scope of risk management to be performed. | Yes | Yes |
| Man.5.BP1a | Define the scope of the strategy and include those life-cycle phases for which the strategy is applicable | No | Yes – ISO 14971 Section 3.4(a) |

Table 5, provides a high level summary of the number of additional sub-practices that needs to be added to the ISO/IEC 15504-5 RM process to achieve the Medi SPICE objectives. The table illustrates 31 sub-practices within Medi SPICE RM. Thirty of these base practices are required within the medical device industry whereas only 8 are required within the ISO/IEC 15504-5 RM process. Only one of the seven ISO/IEC 15504-5 base practices fully satisfies the requirements of the associated Medi SPICE RM practice (i.e BP 4: Analyse Risks). Therefore, the mappings highlight that Medi SPICE needs to be more comprehensive in its coverage of the RM process than ISO/IEC 15504-5 in order to satisfy the regulatory requirements of the medical device industry

Table 5: Summary of Medi SPICE RM Process

| Practice | ISO/IEC 15504-5 Sub-Practices | ISO/IEC 15504-5 Sub-Practices required to meet regulatory medical device requirements | Additional Sub-Practices required to meet regulatory medical device requirements |
| --- | --- | --- | --- |
| BP1: Establish risk management scope | 1 | 1 | 1 |
| BP2: Define risk management strategies | 1 | 1 | 11 |
| BP3: Identify risks | 2 | 1 | 3 |
| BP4: Analyse risks | 1 | 1 | 0 |
| BP5: Define and Perform risk treatment actions | 1 | 1 | 5 |
| BP6: Monitor risks | 1 | 1 | 2 |
| BP7: Take Preventative or Corrective action | 1 | 1 | 1 |
| Total | 8 | 7 | 23 |

With respect to the practices of the Medi SPICE RM process, following the base practices of the ISO/IEC 15504-5 RM process will at best, only partially meet the regulatory requirements of the medical device industry in relation to RM. For RM, the existing ISO/IEC 15504-5 specification of outcomes and base practices can be carried over, with the above mentioned extension into the Medi SPICE framework.

5. Conclusions and Future Work

In the development of  Medi SPICE our approach is to examine all of the ISO/IEC 15504-5 processes focussing initially on the 16 processes that will require the most revision to satisfy the medical device regulatory requirements. Our vision is to provide a framework that will encourage medical device companies to distance themselves from the concept of developing the software first and then completing the necessary documentation that is required to achieve FDA compliance, to instead pursuing a continuous path of software process improvement that will produce more efficient software development and safer medical devices. In this paper we have included a summary of the Medi SPICE RM process as an illustration of a a key process within the medical device industry which requires comprehensive coverage. However, upon development of other Medi SPICE processes we may discover that the associated ISO/IEC 15504-5 base practices will more than satisfy the regulatory medical device requirements and in these cases the medical device industry may be able to adopt some of the additional ISO/IEC 15504-5 practices to increase the efficiency of that process.

6. Acknowledgements

This research is supported by the Science Foundation Ireland (SFI) Stokes Lectureship Programme, grant number 07/SK/I1299.

7. References

- FDA’s Mission Statement - http://www.fda.gov/opacom/morechoices/mission.html. Last accessed June 2009.

- CDRH, General Principles of Software Validation; Final Guidance for Industry and medical device Staff. January 11, 2002

- CDRH, Guidance for the Content of Premarket Submissions for Software Contained in Medical Devices; Guidance for Industry and medical device Staff. May 11, 2005

- CDRH, Off-The-Shelf Software Use in Medical Devices; Guidance for Industry, medical device Reviewers and Compliance. Sept 9, 1999

- Eagles, S. Murray, J. Medical Device Software Standards: Vision and Status, Medical Device and Diagnostic Industry, May 2001, http://www.devicelink.com/mddi/archive/01/05/002.html. Last accessed June 2009.

- ISO/IEC 15504, Information Technology – Process Assessment – Part 5: An exemplar Process Assessment Model, ISO/IEC JTC1/SC7, International Standards Organisation, 2006. http://www.iso.org/iso/iso_catalogue/catalogue_tc/catalogue_detail.htm?csnumber=38934. . Last accessed June 2009.

- European Council, 1993. “Council Directive 93/42/EEC Concerning Medical Devices”, 14 June 1993.

- ANSI/AAMI/ISO 14971, Medical devices – Application of risk management to medical devices, 2nd Edition, 2007

- BS EN 60601-1-4: Medical Electrical Equipment, Part 1 - General requirements for safety. (2000)

- ANSI/AAMI/IEC 62304, Medical device software – Software life cycle processes, Association for the Advancement of Medical Instrumentation, 19-Jul-2006 (replacement for SW68)

- AAMI TIR32:2004, Medical device software risk management, 2005

- ISPE, GAMP Guide for Validation of Automated Systems. GAMP 4, Dec 2001

- GAMP 5, 2008. GAMP 5: International Society for Pharmaceutical Engineering (ISPE): A Risk-Based Approach to Compliant GxP Computerized Systemsan-2008. http://www.techstreet.com/cgi-bin/detail?product_id=1559506. Last accessed June 2009.

- IEC 61508, 2006. IEC 61508 Overview Report, A Summary of the IEC 61508 Standard for Functional Safety of Electrical/Electronic/Programmable Electronic Safety-Related Systems, 2006,  http://www.exida.com/articles/iec61508_overview.pdf. Last accessed June 2009.

- +SAFE, V1.2: A Safety Extension to CMMI-DEV, V1.2, Defence Materiel Organisation, Australian Department of Defence, March 2007,Software Engineering Institute, TECHNICAL NOTE CMU/SEI-2007-TN-006http://www.sei.cmu.edu/pub/documents/07.reports/07tn006.pdf Last accessed June 2009.

- ISO/IEC 15504, Information Technology – Process Assessment – Part 2: Performing an Assessment, ISO/IEC JTC1/SC7, International Standards Organisation, 2003. http://www.iso.org/iso/iso_catalogue/catalogue_tc/catalogue_detail.htm?csnumber=37458 . Last accessed June 2009.

- ISO/IEC 15504 Part 7 - ISO/IEC TR 15504-7:2008, Information technology - Process assessment - Part 7: November 2008. http://www.iso.org/iso/iso_catalogue/catalogue_tc/catalogue_detail.htm?csnumber=50519. Last accessed June 2009.

- IEC 60812, Analysis technique for system reliability - Procedure for failure modes and effects analysis (FMEA), 2006.